## Sheet1
ID | Description | Date of test | Test site | Picture | Material | Process | Creator (contact info and company) | Link to stl and cad files | Test 1: Expert evaluation by 2 ID physicians | Test 2: Collection sufficiency | Test 3: PCR compatibility | Test 4: 24hrs Post autoclave. Bend test | Notes
This is a unique identifier, ideally including the source's name |  |  | E.g., BIDMC | Please take a high-resolution picture of the entire item provided. | Please describe all (known) materials in the bulb head, including plastic, any adhesive, and any flock | Please include a brief description of the design process that an expert would understand well enough to reproduce. For repurposed swabs, write "commercial." | How to contact the creator for questions/further development | Please send .stl and CAD files, which we will post. | If passes, write "Pass". If fails, write "Fail" and give a brief reason why. Should be filled out by someone who has experience actually using NP swabs in a clinical context. Please leave the name of this person. | If a Gram stain of the (inner) cheek made from the swab shows material broadly consistent with our control, then write "Pass". If not, "Fail" and give reason (too abrasive, did not pick up material, etc.) Note: BIDMC protocol does not include adding fluid. | As above, "Pass" vs. "Fail". Please indicate the protocol used. You can post the protocol in the Testing/ directory, and link to it here
Puritan Sterile polyester-tipped applicators 25-806 1PD |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Puritan/puritan_sterile_polyester_tipped_applicator_25-805_1PD.png | Polyester tip | Commercial | Puritan | NA | Did not pass | Not done |  |  | (tip diameter too large)
Fisher polyester-tipped applicators 22363164 |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Fisher/fisher_polyester_tipped_applicators_22363164.png | Polyester tip | Commercial | Fisher | NA | Did not pass | Not done |  |  | (tip diameter too large)
Puritan sterile foam tipped applicator 25-1506 1PF |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Puritan/puritan_sterile_foam_tipped_applicator_25-1506_1PF.png |  | Commercial | Fisher | NA | Pass | Pass | Pass |  | (but tip diameter large, would be difficult to use)
Copan BD ESwab collection and transport system 220532 |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Copan_BD/copan_BD_ESwab_collection_and_transport_system_220532.png |  | Commercial | Fisher | NA | Pass | Pass
Copan ESwab 220245 |  | 2020-03-22 | BIDMC | NA |  | Commercial | Fisher | NA |  |  | Pass
Hologic Aptima LBL-06850 |  | 2020-03-22 | BIDMC | Pending |  | Commercial | Fisher | NA | Pass |  | Pass |  | (but tip diameter large)
Jifei "Carbon" (internal ID, *not* Carbon3D the company) |  | 2020-03-22 | BIDMC | (meth)acrylated urethane photopolymer |  |  | OPT Industries (Jifei Ou) |  |  | Pass | Pass
Jifei Material 2 |  | 2020-03-22 | BIDMC | (meth)acrylated urethane photopolymer | Material 2 |  | OPT Industries (Jifei Ou) |  | NOT SWAB MATERIAL ONLY | NOT SWAB MATERIAL ONLY | Pass
Nilson |  | Mach 20,2020 | BIDMC |  | Rose-print1 | Printed usign E Guide 2.0, cleaned in 99% IPA, 2000 Flash cure with ottoflash unit |  |  |  |  | Pass
Formlabs - Honey Tip |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/FormLabs/honeydipper.png |  |  | FormLabs (Craig Brody) |  | Did not pass | Pass | Pass |  | (shaft needs to be smooth; material too brittle)
Formlabs - Spiral Tip |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/FormLabs/surgical_guide_technical_data_sheet_en.pdf |  |  | FormLabs (Craig Brody) |  | Did not pass | Did not pass (material did not transfer to slide?) | Pass |  | (shaft needs to be smooth; material too brittle)
Formlabs - Bristle Tip |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/FormLabs/bristles.png |  |  | FormLabs (Craig Brody) |  | Did not pass | Did not pass (material did not transfer to slide?) | Pass |  | (shaft needs to be smooth; material too brittle)
Formlabs - Dental LT clear (Bristle, Honey tips) |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/FormLabs/DentalLT.JPG |  |  | FormLabs (Craig Brody) |  | Did not pass |  |  |  | (tip needs to be more flexible - will abutt posterior nasopharynx, but flexibility of NP neck, handle/body still needs to be more smooth)
Formlabs - Dental SG (Fern tip) |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/FormLabs/SG.JPG |  |  | FormLabs (Craig Brody) |  | Did not pass | Pass |  |  | (tip needs to be more flexible - will abutt posterior nasopharynx, but flexibility of NP neck, handle/body still needs to be more smooth)
Formlabs - Grey Pro (Spiral, Honey, Bristle) |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/FormLabs/GreyPro.JPG |  |  | FormLabs (Craig Brody) |  | Did not pass |  |  |  | (material brittle, tip needs to be more flexible - will abutt posterior nasopharynx, but flexibility of NP neck, handle/body still needs to be more smooth)
Formlabs - Durable (Fern, Honey, Bristle) |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/FormLabs/Durable.JPG |  |  | FormLabs (Craig Brody) |  | Did Not Pass | Pass (All but Bristle tip -very few cells with that tip) | Pass |  | *Tested this material as was closest to desirable design although did not pass Test . (Flexibility of tip appropriate, but neck for these designs is too elastic, difficult to pass middle turbinate, and folds in external nasal cavity)
Formlabs Tough (Spiral, Bristle, Honey, Fern) |  | 2020-03-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/FormLabs/tough.JPG |  |  | FormLabs (Craig Brody) |  | Did not pass |  |  |  | (tip needs to be more flexible - will abutt posterior nasopharynx)
Formlabs Dental Model (Honey, Bristle) |  | 2020-03-24 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/FormLabs/model.JPG |  |  | FormLabs (Craig Brody) |  | Did not pass |  |  |  | (tip needs to be more flexible - will abutt posterior nasopharynx, and material too brittle)
Formlabs Tough 1500 (Spiral, Brittle) |  | 2020-03-25 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/FormLabs/tough1500.JPG |  |  | FormLabs (Craig Brody) |  | Did not pass |  |  |  | (tip needs to be more flexible - will abutt posterior nasopharynx)
Formlabs Tough - Tip 4 (Spiral, Bristle, Honey, Fern) |  | 2020-03-26 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/FormLabs/tough.JPG |  |  | FormLabs (Craig Brody) |  | Did not pass |  |  |  | (tip needs to be more flexible - will abutt posterior nasopharynx)
Harvard Army |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Harvard_army/new_skin.png |  |  | Army (Kit Parker) |  | Did not pass | "New Skin" passes; others did not pass (too little material) |  |  | (3 glues used - Locktite, Dermabond, New skin) (material too brittle)
Harvard Army Ultem |  | 2020-03-23 |  | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Harvard_army/ultem.JPG |  |  | Army (Kit Parker) |  | Did not pass |  |  |  | (material too brittle, tip inflexible
Harvard Army Onyx |  | 2020-03-23 |  | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Harvard_army/onyx.JPG |  |  | Army (Kit Parker) |  | Did not pass |  |  |  | (tip too inflexible)
Harvard Army Polyjet |  | 2020-03-23 |  |  |  |  | Army (Kit Parker) |  | Pass | (not tested, only one swab which was tested for PCR) | Pass
W1 |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/wyss_w1.png | rayon, fiber, silicon glue, polycarbonate probe |  | Wyss Institute (Adama Sesay) |  | Pass | Pass | Pass
W2 |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/wyss_w2.png | nylon fiber (0.75µm), silicon glue, polycarbonate probe |  | Wyss Institute (Adama Sesay) |  | Did not pass | Did not pass (No visible cells or bacteria) |  |  | (needs a tapered, flexible end)
W3 |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/wyss_w3.png | rayon fiber, silicon glue, nylon probe |  | Wyss Institute (Adama Sesay) |  | Pass | Did not pass (very few cells/bacteria)
W4 |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/wyss_w4.png | rayon immfiber, silicon glue, commercial straw |  | Wyss Institute (Adama Sesay) |  | Did not pass | Pass |  |  | (needs a tapered, flexible end)
W5 |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/wyss_w5.png | nylon fiber (0.75µm), nylon probe, silicon glue |  | Wyss Institute (Adama Sesay) |  | Pass | Pass
W6 |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/wyss_w6.png | polycarbonate 3D-printed probe |  | Wyss Institute (Adama Sesay) |  | Pass | Pass
W7 |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/wyss_w7.png | nylon 3D-printed probe |  | Wyss Institute (Adama Sesay) |  | Pass | Pass
W8 |  | 2020-03-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/wyss_w8.png | nylon fiber (0.75µm), silicon glue, polycarbonate probe |  | Wyss Institute (Adama Sesay) |  | Pass | Pass
VWR 1ul loop + pellon p45 fabric knot - Jenny |  | 2020-03-22 | HH (Seattle) | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Jenny/DIYswab/ | polystyrene loop and pellon P45 fabric | see https://github.com/HMS-RIC/Covid19-NP-Swab/wiki/DIY-swab-assembly | Jenny |  | Pass | need review of smear pics |  |  | (doctor/swabber was able to use 5mm loop, newer iteration is <3mm) - Jenny
USF_1_1 | Rings w/ medium spacing distance | 2020-03-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/USF/USF1-1.jpg |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass | Pass | Pass |  | (tip is too sharp but appropriately flexible, but body design (flexibility of neck and stiffness of handle is appropriate). There is a fine powder on these materials that we are concerned could be health hazard if aspirated
USF_1_2 | Rings w/ low spacing distance |  |  | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/USF/USF1-2.JPG |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass | Pass | (presume pass if same material as USF1_1) |  | (tip is too sharp but stiff/brittle, but body design (flexibility of neck and stiffness of handle is appropriate). There is a fine powder on these materials that we are concerned could be health hazard if aspirated
USF_1_3 | Rounded tip w/ protrusions |  |  | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/USF/USF1-3.JPG |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass | Pass | (presume pass if same material as USF1_1) |  | (tip is too sharp and stiff, but body design (flexibility of neck and stiffness of handle is appropriate). There is a fine powder on these materials that we are concerned could be health hazard if aspirated
USF_1_4 | Screw; finely threaded |  |  | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/USF/USF1-4.JPG |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass | Pass | (presume pass if same material as USF1_1) |  | (tip is too sharp and stiff/brittle, but body design (flexibility of neck and stiffness of handle is appropriate). There is a fine powder on these materials that we are concerned could be health hazard if aspirated
USF_1_5 | Screw; medium thread |  |  | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/USF/USF1-5.JPG |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass | Pass | (presume pass if same material as USF1_1) |  | (tip is too sharp but we like the spring and how flexible it is, but body design (flexibility of neck and stiffness of handle is appropriate). There is a fine powder on these materials that we are concerned could be health hazard if aspirated
USF_1_6 | Screw;large thread | 2020-03-23 |  | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/USF/USF1-6.JPG |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass | Pass | (presume pass if same material as USF1_1) |  | (tip is too sharp but we like the spring and how flexible it is, but body design (flexibility of neck and stiffness of handle is appropriate). There is a fine powder on these materials that we are concerned could be health hazard if aspirated
HP_H | HP swabs with holes | March 23,2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HPH.JPG |  |  | HP (LihuA Zhao; Zhao, Lihua <lihua.zhao@hp.com) |  | **Conflicting Results** Pass per Jenny Hu 032320 Not pass per BIDMC: | **Conflicting Results** Pass per Jenny Hu 032320 (collected a few cells/bacteria from cheek swab then rubbing directly on slides; better results from rinsing the tip with 150ul of buffer then transfering liquid to slide). Not pass per BIDMC protocol (no elution); very few cells collected |  |  | Doing some more mechanical measurements/testing this afternoon then my doctor friend will stop by after work and I'll do NP swabs with Copan, HP_H, HP_NH, VWRloop+facric. Will smear swabs directly on slides then try rinsing and looking at rinse liquid for cells/bacteria. per Jenny Hu 032320: doctor swabber would like the shaft to be longer but likes the shaft flexibility and tip size. : tip not very flexible and too sharp at the end
HP_NH | HP swabs without holes | 2020-03-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HPNP.JPG |  |  | HP (LihuA Zhao; Zhao, Lihua <lihua.zhao@hp.com) |  | **Conflicting Results** Pass per Jenny Hu 032320:. Not pass per BIDMC: | **Conflicting Results** Pass per Jenny Hu 032320 (collected a few cells/bacteria from cheek swab then rubbing directly on slides; better results from rinsing the tip with 150ul of buffer then transfering liquid to slide). Not pass per BIDMC protocol (no elution); very few cells collected | Pass |  | doctor swabber would like the shaft to be longer but likes the shaft flexibility and tip size. tip not very flexible and too sharp at the end
C_1 | Carbon 3D bag 1 | 2020-03-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C1.JPG |  |  | Carbon 3D ( Hardik Kabaria <hardik@carbon3d.com) |  | Pass | Pass | Pass |  | (good stiffness of handle, and flexible tip), tapered more flexible neck would be better
C_2 | Carbon 3D bag 2 | 2020-03-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C2.JPG |  |  | Carbon 3D ( Hardik Kabaria <hardik@carbon3d.com) |  | Pass | Pass |  |  | (good stiffness of material, and flexible tip), tapered more flexible neck would be better
P_0 | Swab_test_Wyss | 2020-03-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Gorelik/P0.JPG | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Did not pass | Pass |  |  | - tip needs to be more flexible, but stiffness of neck and elasticity of material appropriate)
P_1 | Solid_Circular | 2020-03-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Gorelik/P1.JPG | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Did not pass | Pass |  |  | (tip too sharp, appropriately flexible, but would make tip end less jagged)
P_2 | Spring_Circular | 2020-03-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Gorelik/P2.JPG | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Did not pass | Pass | Pass |  | (tip too sharp but more appropriately flexible, but would make tip end less jagged)
P_3 | Spring_Flat | 2020-03-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Gorelik/P3.JPG | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Pass - | Pass |  |  | tip is softer/flexible, and stiffness of neck and elasticity of material appropriate)
P_4 | Solid_Flat | 2020-03-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Gorelik/P4.JPG | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Did not pass | Pass |  |  | (tip too sharp/stiff, body design appropriate)
P_5 | Swab_Concentric | 2020-03-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Gorelik/P5.JPG | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Pass | Pass |  |  | (tip more appropriately flexible, but would make tip end less jagged)
P_6 | Swab_concentric_bigger_pitch | 2020-03-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Gorelik/P6.JPG | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Pass | Pass |  |  | (tip more appropriately flexible, but would make tip end less jagged)
P_7 | Swab_test_head_4_wyss | 2020-03-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Gorelik/P7.JPG | PA12 |  | Empire Group. (Pavel Gorelik; oreshki@gmail.com ) |  | Pass for tip head
J_1 | flexible swab w/ case | March 24th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/OPT_industries/J_1.JPG | methylacrylic urethane |  | opticalindustries, jifei jifei@optindustries.com |  | Pass but design too short | Pass | passed
USF_2_1 | swab w/ fine nubs | March 24th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/USF/USF_2_1.JPG |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass | Pass | Pass |  | (tip too sharp and should be more flexible, but body design is appropriate)
USF_2_2 | swab w/ finely spaced rings | Marh 24th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/USF/USF_2_2.JPG |  |  | HP (Rick Fulop; ricfulop@desktopmetal.com) |  | Did not pass | Pass |  |  | (tip too sharp and but apporpirately flexible, but body design is appropriate)
B_1 | swab w/very fine nubs | March 24th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Baker/B_1.JPG |  |  | Daniel Baker |  | Did not pass | Pass | Pass |  | (tip is too sharp and ideally more flexible tip, but body design is appropriate)
S_1 | Swab w/ thick thread | March 24th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Seriano/S_1.JPG |  |  | David Seriano |  | Did not pass | Pass | Pass |  | (tip is too sharp, but apporpiately flexible and body design is appopriate)
S_2 | swab w/ fine spring | March 24th | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Seriano/S_2.JPG |  |  | David Seriano |  | Did not pass | Pass |  |  | (tip is too sharp, but very flexible, perhaps too flexible, S1 may be better, and body design is appopriate)
S_3 | swab w/ fine thread | March 24th | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Seriano/S_3.jpg |  |  | David Seriano |  | Did not pass | Pass |  |  | (tip is too sharp, and body design is appopriate)
S_4 | swab w/finely speaced rings | March 24th | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Seriano/S_4.JPG |  |  | David Seriano |  | Did not pass | Pass |  |  | (tip is too sharp, and body design is appopriate, best design of the S-series in terms of head softness)
S_5 | swab w/ low ring spacing | March 24th | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Seriano/S_5.JPG |  |  | David Seriano |  | Did not pass | Pass |  |  | (tip is too sharp, and body design is appopriate)
S_6 | swab w/fine rings | March 24th | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Seriano/S_6.JPG |  |  | David Seriano |  |  | Pass
FT_1 | Foam Swab. Pt: HT1171B | March 24th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Foamtec/FT_1.JPG |  |  | Foamtec International; Michael Strauss |  | Did not pass | Did not pass (not many cells collected) |  |  | (diameter of tip too large, flexible bendable neck would be better)
FT_2 | Foam Swab. Pt: HT1501T2-5 | March 24th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Foamtec/FT_2.JPG |  |  | Foamtec International; Michael Strauss |  | Did not pass | Did not pass (not many cells collected) |  |  | (diameter of tip too large, flexible bendable neck would be better)
FT_3 | Foam. Pt. HT1501TC-5 | March 24th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Foamtec/FT_3.JPG |  |  | Foamtec International; Michael Strauss |  | Did not pass | Did not pass (not many cells collected) |  |  | (diameter of tip too large, flexible bendable neck would be better)
FT_4 | Foam Swab. Pt: HT1001-5 | March 24th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC//img/Foamtec/FT_4.JPG |  |  | Foamtec International; Michael Strauss |  | Did not pass | Did not pass (not many cells collected) |  |  | (diameter tip, diameter of tip too large, flexible bendable neck would be better)
EV_1 | DOWEL. For material testing | March 24th, 2020 | BIDMC |  | PCA 2000. 10 minutes. 100% |  | EnvisionTEC. Al Siblani. asiblani@envisiontec.com |  |  |  | passed
EV_2 | DOWEL. For material testing | March 24th, 2020 | BIDMC |  | PCA100 w/nitrogen | 2x10seconds w/nitrogen | EnvisionTEC. Al Siblani. asiblani@envisiontec.com |  |  |  | off for testing
EV_3 | DOWEL. For material testing | March 24th, 2020 | BIDMC |  | PCA 100 w/nitrogen | 2x30seconds UV w/nitrogen | EnvisionTEC. Al Siblani. asiblani@envisiontec.com |  |  |  | off for testing
EV_4 | DOWEL. For material testing | March 24th, 2020 | BIDMC |  | PCA 100 w/nitrogen | 2x20seconds w/nitrogen | EnvisionTEC. Al Siblani. asiblani@envisiontec.com |  |  |  | off for testing
PVA | Chemical for material testing. Cas: 9002-89-5 | March 25th, 2020 | BIDMC |  | Polly(Vinyl Alcohol) |  | Sigma-Aldrich. Adama; Wyss |  |  |  | passed
Fil_Flt | White filament (for material testing) | March 25h, 2020 | BIDMC |  | 3DF-2 (Raw) (Polyurethane + PVA) |  | Adama; Wyss |  |  |  | passed
Fil_Lay | White filament (for material testing) | March 25th, 2020 | BIDMC |  | 3DF-1 (polyurethane + PVA) |  | Adama Sesay, Wyss Institude |  |  |  | passed
Sc_M | Jar filled with white substance | March 25th, 2020 | BIDMC |  | Scotch Magic Tape |  | 3M. dropped off by Adama Sesay; Wyss Institute |  |  |  | passed
Sc_CC | Jar filled with white substance | March 25th, 2020 | BIDMC |  | Scotch Crystal Clear |  | 3M. dropped off by Adama Sesay; Wyss Institute
W_2_1 | White swab w/ silicon glue and sugar textured head | March 25th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/W_2_1.jpg | SG1. Silicon glue and sugar |  | Wyss Institute (Adama Shesay) |  | Did not pass | Pass |  |  | (body design is appropriate, but tip/neck is too flexible and seems would not easily pass beyond external nares, seems non-uniform & wasn't sure if there would be scale-up reproducibility issues)
W_2_2 | Black swab w/ silicon glue and sugar textured head | March 25th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/W_2_2.jpg | SG2. Silicon glue and Sugar |  | Wyss Institute (Adama Shesay) |  | Pass | Pass |  |  | (body design is appropriate, the tip end material seems very absorbent which we like, but the tip would ideally be more flexible, tip seems non-uniform & wasn't sure if there would be scale-up reproducibility issues)
W_2_3 | Rayon and crystal clear glue textured head | March 25th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/W_2_3.jpg | 3M1. Rayon fiber and scotch crystal clear glue head with polycarbonate swab |  | Wyss Institute (Adama Shesay) |  | Did not pass | Pass |  |  | (tip sharp and inflexible but body design is appropriate, tip seems non-uniform & wasn't sure if there would be scale-up reproducibility issues)
W_2_4 | Rayon and magic tape textured head | March 25th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/W_2_4.jpg | 3M2. Rayon fiber with scoth magic tape swab head on a nylon swab |  | Wyss Institute (Adama Shesay) |  | Pass | Did not do cheek swab, uncured glue |  |  | (tip is flexible and body design is appropriate)
W_2_5 | Rayon fiber and scotch crysal clear textured head | March 25th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/W_2_5.jpg | 3M3 Rayon Fiber with scotch crystal clear textured head on nylon swab |  | Wyss Institute (Adama Shesay) |  | Pass | Pass |  |  | (tip is flexible and body design is appropriate, tip seems non-uniform & wasn't sure if there would be scale-up reproducibility issues, also seems stikcy)
W_2_6 | Rayon fiber/ Scotch magnetic tape textured head | March 25th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/W_2_6.jpg | 3M4 Rayon Fiber/Scotch magic tape textured head on polycarbonate head |  | Wyss Institute (Adama Shesay) |  | Did not pass | Did not do cheek swab, uncured glue |  |  | (tip is sharp and inflexible but body design is appropriate, tip seems non-uniform & wasn't sure if there would be scale-up reproducibility issues)
W_2_7 | Nylon fiber/magic tape | March 25th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/W_2_7.jpg | 3M5 Nylon Fiber/Magic Tape textured head on polycarbornate swab |  | Wyss Institute (Adama Shesay) |  | Pass | Pass |  |  | (tip would ideally be more flexible but body design is appropriate, tip seems non-uniform & wasn't sure if there would be scale-up reproducibility issues)
W_2_8 | Nylon fiber/magic tape | March 25th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/W_2_8.jpg | 3M6 Nylon Fiber/magi tape textured head on nylon swab |  | Wyss Institute (Adama Shesay) |  | Pass | Pass |  |  | (tip ideally would be more flexible, but design is appopriate, tip seems non-uniform & wasn't sure if there would be scale-up reproducibility issues)
W_2_9 | Nylon FIber/Crystal Clear | March 25th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/W_2_9.jpg | 3M_7 Nylon fiber/ crystal clear textured head on a nylon swab |  | Wyss Institute (Adama Shesay) |  | Pass | Pass |  |  | (tip is flexible and body design is appropriate, tip seems non-uniform & wasn't sure if there would be scale-up reproducibility issues)
W_2_10 | Nylon fiber/magic tape | Martch 25th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/W_2_10.jpg | 3M8 nylon fiber/magic tape texted head on a nylon swab |  | Wyss Institute (Adama Shesay) |  | Pass | Pass |  |  | (tip is more flexible which we like but body design is appropriate, tip seems non-uniform & wasn't sure if there would be scale-up reproducibility issues)
W_2_11 | Nylon FIber/ crytal clear textured head | March 25th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/W_2_11.jpg | 3M9 nylon fiber w/ crystal clear textured head on polycarb |  | Wyss Institute (Adama Shesay) |  | Did not pass | Pass |  |  | (too sharp and tip too inflexible)
W_2_12 | Sugar on nylon | March 25th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/W_2_12.jpg | PDMS [sugar on nylon] |  | Wyss Institute (Adama Shesay) |  | Did not pass | Pass | Pass |  | (tip/neck are too flexible we feel like it would get caught in the external nares and not pass deeper well)
EV_2_1 | 3cm rods (for material testing) |  |  | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Envisiontec/EV_2_1.jpg | C 29C |  | EnvisionTEC. Al Siblani. asiblani@envisiontec.com |  | NA |  | passed
C_3 | 4mm voro | March 26th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C_3.jpg | keysplint Soft Clear | dreve PCU 90 N2 15 min | Carbon3D. Keshav Swarup |  | Did not pass | Pass | *pass (presumably same material as C!) |  | (tip head too firm, would create a neck for more flexibility at the end, and would put breakpoint from tip head end at least 7-10cm from the end.
C_4 | 3mm voro | March 26th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C_4.jpg | keysplint Soft Clear | dreve PCU 90 N2 15 min | Carbon3D. Keshav Swarup |  | Did not pass | Pass | *pass (presumably same material as C!) |  | (tip head too firm, would create a neck for more flexibility at the end, and would put breakpoint from tip head end at least 7-10cm from the end).
C_5 | USF | March 26th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C_5.jpg | keysplint Soft Clear | dreve PCU 90 N2 15 min | Carbon3D. Keshav Swarup |  | Did not pass | Did not pass | *pass (presumably same material as C!) |  | (tip head too sharp and stiff although bristeled head may perform better ultmately for specimen capture. Body design is good)
C_6 | long bulb | March 26th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C_6.jpg | keysplint Soft Clear | dreve PCU 90 N2 15 min | Carbon3D. Keshav Swarup |  | Pass | Pass | *pass (presumably same material as C!) |  | **likely a good design, the head is longer and flexible and probably can pass deeper as well as pickup more material. (unique design, and may be preferable to others given logner head/more absorption, and flexibility of entire head while maintaining shape well after deformation)
C_7 | 3mm 1.5mm neck | March 26th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C_7.jpg | keysplint Soft Clear | dreve PCU 90 N2 15 min | Carbon3D. Keshav Swarup |  | Pass | Pass | *pass (presumably same material as C!) |  | (body design std Copan, tip soft enough)
HP_FiT | finely threaded black swab | March 27th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_Fit.jpg | PA12 | 580 platform | HP. Isabel Sanz (isabel.sanz-ananos@hp.com) |  | Did not pass | Did not pass. Very few cells collected |  |  | (tip head is too frim)
HP_nub | fine nubs on black swab | March 27th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_nub.jpg | PA12 | 580 platform | HP. Isabel Sanz (isabel.sanz-ananos@hp.com) |  | Did not pass | Pass |  |  | (tip head is too frim)
HP_WT | black swab with wide threads | March 27th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_WT.jpg | PA12 | 580 platform | HP. Isabel Sanz (isabel.sanz-ananos@hp.com) |  | Did not pass | Did not pass. No cells collected
HP_WS | wide screw tiped black swab | March 27th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_WS.jpg | PA12 | 580 platform | HP. Isabel Sanz (isabel.sanz-ananos@hp.com) |  | Pass | Pass | Pass |  | *post-autoclave the material feels and performs the same. (ideally tip point would be less sharp, although material is much preferable in this verison than prior HP designs)
HP_FS | fine spring tiped black swab | March 27th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_FS.jpg | PA12 | 580 platform | HP. Isabel Sanz (isabel.sanz-ananos@hp.com) |  | DId not pass | Pass |  |  | (tip head is too flexible, we are concerned could get caught in external nares, and unable to reach posteiror nasopharynx due to such high flexibility)
C_2_1 | long hexagon tip with neck | March 27th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C_2_1.jpg | Keysplint Soft Clear | Dreve PCU 90 N2 20 minutes | Carbon 3D ( Hardik Kabaria <hardik@carbon3d.com) |  | Pass | Pass |  |  | (but tip point should be less sharp, although material is much preferable in this verison than prior HP designs)
HA_2_Onyx | black swab with square wave pattern | March 27th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Harvard_army/HA_2_Onyx.jpg |  |  | Daniel Davis |  | Did not pass |  |  |  | (tip too sharp)
HA_V13 | yelow swab w/ bulb | March 27th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Harvard_army/HA_V13.jpg | PEI |  | Daniel Davis |  | Did not pass |  |  |  | (sharp points at breakoff points)
HA_2_P | yellow with square wave | March 27th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Harvard_army/HA_2_P.jpg | PEI |  | Daniel Davis |  | Did not pass |  |  |  | (sharp points at breakoff points)
HA_D5 | #5 Dymax md-1405-T-ur-4r-sc | March 27th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Harvard_army/HA_D5.jpg | PEI |  | Daniel Davis |  | Did not pass |  | Pass |  | (would have a flexible neck to allow further bending in the posterior nasopharynx)
HA_D6 | #6 Dymax md 215-CTH-SV01-WR-SC | March 27th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Harvard_army/HA_D6.jpg | PEI |  | Daniel Davis |  | Did not pass |  |  |  | (would have a flexible neck to allow further bending in the posterior nasopharynx)
HA_D7 | #7 Dymax md 1072-m | March 27th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Harvard_army/HA_D7.jpg | PEI |  | Daniel Davis |  | Did not pass | Did not test, only one rec'd |  |  | (would have a flexible neck to allow further bending in the posterior nasopharynx)
HA_D8 | #8 Dymax 1120-m-ur | March 27th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Harvard_army/HA_D8.jpg | PEI |  | Daniel Davis |  | Did not pass | Did not test, only one rec'd |  |  | (would have a flexible neck to allow further bending in the posterior nasopharynx)
HA_D9 | #9 3M Super 77 | March 27th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Harvard_army/HA_D9.jpg | PEI |  | Daniel Davis |  | Did not pass | Did not test, only one rec'd | Pass |  | (would have a flexible neck to allow further bending in the posterior nasopharynx)
HA_D10 | #10 water based vinyl acetate emulsion, cured 160C | March 27th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Harvard_army/HA_D10.jpg | ? |  | Daniel Davis |  | Did not pass |  | Pass |  | (would have a flexible neck to allow further bending in the posterior nasopharynx)
EV_3_1 | Envision One printed Swab Model Version 11 | March 30th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/EnvisionTec/EV_3_1.jpg | Material E Guide | IPA wash 2x Cured 10 min PCA 2000 Code, Rose-Print -2 | EnvisionTEC. Al Siblani. asiblani@envisiontec.com |  | Did not pass | Pass |  |  | material seems similar post-autoclave, but is still sharp and somewhat brittle. - tip seems sharp and ideally should be more flexible, but body design seems appropriate
EV_3_2 | Envision One printed Swab Model Version 11--altered break point |  | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/EnvisionTec/EV_3_2.jpg | Material E Guide | IPA wash 2x Cured 10 min PCA 2000 Code, Rose-Print -2 | EnvisionTEC. Al Siblani. asiblani@envisiontec.com |  | Pass | pass |  |  | but the tip is somewhat brittle/stiff, would be better to be slightly more flexible and would be concerned about possible breaking in the nasopharynx
MF_1 | black swab with floc | March 30th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Markforged/MF_1.jpg | Nylon reinforced with carbon fiber |  | Markforged. Dan Eiref.daniel.eiref@markforged.com |  | Did not pass | Pass |  |  | body design is appropriatley flexible and stiff but tip diameter head seems large and would be difficult to pass to posterior nasopharynx
T_S | gray swab with blunted screw | March 30th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Tooker/T_S.jpg | ?HP |  | Angela Tooker |  | Did not pass | Passed |  |  | tip sides are sharp, although tip end seems appropriatley blunted, material seems somewhat powdery and unfinished
T_BS | gray swab with bluned spring | March 30th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Tooker/T_BS.jpg | ?HP |  | Angela Tooker |  | Did not pass | Passed |  |  | tip sides are sharp, although tip end seems appropriatley blunted and flexible, material seems somewhat powdery and unfinished
T_SC | gray swab with multiple square cavities | March 30th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Tooker/T_SC.jpg | ?HP |  | Angela Tooker |  | Did not pass | Did not pass | Pass |  | tip sides are more smooth but still overall tip head seems sharp although appropriatley blunted, material seems somewhat powdery and unfinished
MF_2 | White swab with flocked loop | March 30th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img//Markforged/MF_2.jpg | medical grade nylon with flocked tip |  | Markforged. Tom Allen;Greg Mark |  | Pass | Passed | Pass |  | body design appropriately flexible and stiff, but would need a breakpoint
J_2 | Reinforced swab with larger distal tip diameter | March 30th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/OPT_industries/J_2.JPG | meth(acrylated urethane photopolymer |  | OPT Industries (Jifei Ou) |  | Did not pass | Passed |  |  | body design is too flexible, needs to be stiffer in order to appropriately maintain shape to pass to posterior nasopharynx, also material seems tacky/sticky to touch, design of tip seems absorbent and appropriatley flexible
J_3 | Reinforced swab with mesh tip | MARch 30th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/OPT_industries/J_3.JPG | meth(acrylated urethane photopolymer |  | OPT Industries (Jifei Ou) |  | Did not pass | Passed |  |  | body design is too flexible, needs to be stiffer in order to appropriately maintain shape to pass to posterior nasopharynx, also material seems tacky/sticky to touch, design of tip seems absorbent and appropriatley flexible
TS_F_6 | foam swab, 6inch | March 30th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Teknipure/TS_F_6.jpg | TS_F_6 |  | Teknipure. Chris Heiland |  | Did not pass | Passed | Passed |  | tip diameter head on the larger side as well as body design would preferably have a narrow-diameter neck ~3.5cm which will be easier to manipulate and pass through to the posterior nasopharynx
TS_MW_6 | foam swab, 6inch | March 30th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Teknipure/TS_MW_6.jpg | TS_MW_6 |  | Teknipure. Chris Heiland |  | Did not pass | Passed |  |  | tip diameter head on the larger side as well as body design would preferably have a narrow-diameter neck ~3.5cm which will be easier to manipulate and pass through to the posterior nasopharynx
TS_FMD_6 | foam swab, 6inch | March 30th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Teknipure/TS_FMD_6.jpg | TS_FMD6 |  | Teknipure. Chris Heiland |  | Did not pass | Did not pass |  |  | tip diameter head on the larger side as well as body design would preferably have a narrow-diameter neck ~3.5cm which will be easier to manipulate and pass through to the posterior nasopharynx
TS_P_6 | foam swab, 6inch | March 30th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Teknipure/TS_P_6.jpg | TS_P_6 |  | Teknipure. Chris Heiland |  | Did not pass | Did not pass |  |  | tip diameter head on the larger side as well as body design would preferably have a narrow-diameter neck ~3.5cm which will be easier to manipulate and pass through to the posterior nasopharynx
E_3_3 | EV swab with ?altered tip head | March 30th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/EnvisionTec/EV_3_3.jpg | Material E Guide |  | EnvisionTEC. Al Siblani. asiblani@envisiontec.com |  | Pass | Passed |  |  | but the tip is somewhat brittle/stiff, would be better to be slightly more flexible and would be concerned about possible breaking in the nasopharynx
BE_1465.10 | foam swab, LTO1465 | March 30th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Berkshire_Engineered_Clean/BE_1465.10.jpg | LTO1465.10 |  | Berkshire Engineered clean. |  | Did not pass | Did not pass |  |  | tip diameter head on the larger side as well as body design would preferably have a narrow-diameter neck ~3.5cm which will be easier to manipulate and pass through to the posterior nasopharynx
O_2A | Origin.IO swab design 2A | March 30th, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Origin/O_2A.jpg | ? |  | Origin. Stuart Steinbock (stuart@origin.io) |  | Pass | Passed | Pass |  | but tip head is large in diameter, and ideally would be more softer/flexible, would likely be very uncomfortable to pass into posterior nasopharynx
O_2A_1 | Origin.IO swab desing with hexagonal mesh | March 31st, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Origin/O_2A_1.jpg | ? |  | Origin. Stuart Steinbock (stuart@origin.io) |  | Pass | Pass | (same material as O_2A) |  | *passed after autoclave, material remains similar to original, we did like the O_2A_2 head better. but tip head on 2A_2 preferbale
O_2A_2 | Origin.IO swab desing with small hexagonal mesh | March 31st, 2020 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Origin/O_2A_2.jpg | ? |  | Origin. Stuart Steinbock (stuart@origin.io) |  | Pass | Pass | (same material as O_2A) |  | (body design is appropriate, and tip head is softer)
SB-71-4582 | injection molded w/ foam sock | 3/31/2020, | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/SuperBrush/SB_71_4582.jpg | 71-4582 |  | Johan Sunryd jsunryd@superbrush.com |  | Did not pass | Pass |  |  | (body design would ideally have a more flexible neck as in design 52-1005 so that swab can be pushed into posterior nasopharynx)
SB_52-1005 | injection molded w/foam sock | 2020-03-31 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/SuperBrush/SB_52_1005.jpg | 52-1005 |  | Johan Sunryd jsunryd@superbrush.com |  | Pass | Pass |  |  | (tip head with appropriate diameter and body design is appropriate, body design needs an easier breakpoint however (preferably ~7cm from end of tip head)
C_3_1 | 100 non sterile, single packed swabs | 2020-03-31 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C_3_1.jpg | Keysplint Soft Clear | Dreve PCU 90 N2 20 minutes | Hardik Kabaria hardik@carbon3d.com | [5 empty] | *post-autoclave, this material feels and performs similraly
J_2_2 | Jiffei no.2 | 2020-03-31 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/OPT_industries/J_2_2.jpg | meth(acrylated urethane photopolymer | ? | opticalindustries, jifei jifei@optindustries.com |  | Did not pass
J_2_3 | Jiffei no.3 | 2020-03-31 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/OPT_industries/J_2_3.jpg | meth(acrylated urethane photopolymer | ? | opticalindustries, jifei jifei@optindustries.com |  | Did not pass |  |  |  | material is brittle although more stiff and somewhat rough
J_2_4 | Jiffei no.4 | 2020-03-31 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/OPT_industries/J_2_4.jpg | meth(acrylated urethane photopolymer | ? | opticalindustries, jifei jifei@optindustries.com |  | Did not pass |  |  |  | appropriate body design and stiffness, but material is still brittle, and somewhat rough
J_2_5 | Jiffei no.5 | 2020-03-31 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/OPT_industries/J_2_5.jpg | meth(acrylated urethane photopolymer | ? | opticalindustries, jifei jifei@optindustries.com |  | Did not pass |  |  |  | appropriate body design and stiffness, but material is still brittle and somewhat rough
HP_Spring_gray | HP gray swab with cylinder capped spring tip | 2020-04-01 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_Spring_gray.jpg | PA12 | 580 platform | HP (LihuA Zhao; Zhao, Lihua <lihua.zhao@hp.com) |  | Did not pass |  | (same as HP_ws) |  | polished design better, and feels smoother, breakpoint is a lot easier than the non-polished
HP_Spiral_design | HP gray spiral design with hemisphere tip | 2020-04-01 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_Spiral_Design.jpg | PA12 | 580 platform | HP (LihuA Zhao; Zhao, Lihua <lihua.zhao@hp.com) |  | Did not pass |  | (same as HP_ws) |  | polished design better, and feels smoother, breakpoint is a lot easier than the non-polished
HP_spring_polished | HP black spring design with cylinder cap | 2020-04-01 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_Spring_Polished.jpg | PA12 | 580 platform | HP (LihuA Zhao; Zhao, Lihua <lihua.zhao@hp.com) |  | Pass | Pass | (same as HP_ws)
HP_spiral_polished | HP black spiral design with hemisphere cap | 2020-04-01 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_Spiral_Polished.jpg | PA12 | 580 platform | HP (LihuA Zhao; Zhao, Lihua <lihua.zhao@hp.com) |  | Pass | Pass | (same as HP_ws)
EV_3_1 Autoclaved 2020-03-31 | EV_3_1 Autoclaved 2020-03-31 | 2020-04-01 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/EnvisionTec/EV_3_1.jpg | Material E Guide | IPA wash 2x Cured 10 min PCA 2000 Code, Rose-Print -2. [Steam 270/4/30inch] | [5 empty] | Passed
O_2A_1 Autoclaved 2020-03-31 | O_2A_1 Autoclaved 2020-03-31 | 2020-04-01 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Origin/O_2A_1.jpg | ? | ? [Steam 270/4/30inch] | [5 empty] | Passed
C_3_1 Autoclaved 2020-03-31 | C_3_1 Autoclaved 2020-03-31 | 2020-04-01 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C_3_1.jpg | Keysplint Soft Clear | Dreve PCU 90 N2 20 minutes | [5 empty] | Passed
HP_WS Autoclaved 2020-03-31 | HP_WS Autoclaved 2020-03-31 | 2020-04-01 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_WS.jpg | PA12 | 580 platform |  |  |  |  | (same as HP_ws) | Passed
HP_Spring_Polished_1 | HP_Spring_Polished with modified breakoff point | 2020-04-02 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_Spring_Polished_1.jpg | PA12 | 580 platform
HP_Spiral_Polished_1 | HP_Spiral_Polished with modified breakoff point | 2020-04-02 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_Spiral_Polished_1.jpg | PA12 | 580 platform
MF_2_Br | MF_2 with breakoff point | 2020-04-02 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Markforged/MF_2_Br.jpg | Medical nylon | Markforged. Tom Allen;Greg Mark | [5 empty] | passed
WIM-001 | Wyss swab w/ rings | 2020-04-02 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/WIM_001.jpg | Polypropylene |  | Richard Novak. Richard.Novak@wyss.harvard.edu |  | Pass | Pass. Much less material than WIM-002. |  | Evaluator 1: appropriate body design, and tip head appropriately flexible and stiff. Body design should be smoother (some support attachments points too sharp) and breakpoint should be easier. Evaluator 2: tip is sharp and inflexible
WIM-002 | Wyss swab w/ rings | 2020-04-02 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/WIM_002.jpg | nylon |  | Richard Novak. Richard.Novak@wyss.harvard.edu |  | Pass | Pass, but not much material. |  | Evaluator 1: but tip head preferable from 001. Body design should be smoother (some support attachments points too sharp) and breakpoint should be easier. Evaluator 2: tip is sharp and inflexible
O_2A_2_1 | O_2A_2 tip on O_2A_1 neck | 2020-04-03 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Origin/O_2A_2_1.jpg | ? | ? |  |  | pass | Pass
O_2A_2_2 | O_2A_2_2 new tip for smoothness | 2020-04-03 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Origin/O_2A_2_1.jpg | ? | ? |  |  | pass | Pass |  | Passed
EV_2_2 | Envisiontec autoclaved swab | 2020-04-03 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/EnvisionTec/EV_2_2.jpg | Material E Guide Soft | IPA wash. Cured 15 min 60 PCA 2000, Rose-Print-S. Sterilized 270F in an M11 Ultraclave | EnvisionTEC. Al Siblani. asiblani@envisiontec.com |  | Pass | Pass |  | -edges of tip are sharp
EV_2_3 | Envisiontec autoclaved swab with new head | 2020-04-05 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/EnvisionTec/EV_3_3.jpg | Material E Guide Soft | IPA wash. Cured 15 min 60 PCA 2000, Rose-Print-S. Sterilized 270F in an M11 Ultraclave | EnvisionTEC. Al Siblani. asiblani@envisiontec.com |  | pass
HP_WS_DG | HP_WS with hemisphere tip on a dark gray swab | 2020-04-06 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_WS_DG.jpg | PA12 | 580 platform | HP (LihuA Zhao; Zhao, Lihua <lihua.zhao@hp.com) |  | Pass | pass |  | rounded tip makes the swab much more comfortable, the finish seems a little rough and powdery. Acid-vapor-wash treatment seems to resolve this issue
MF_2_Br2 | Markforged swab with new breakpoint placement | 2020-04-06 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Markforged/MF_2_Br2.jpg | medical nylon |  | Markforged. Dan Eiref.daniel.eiref@markforged.com |  | DId not pass |  |  | breakpoint at neck is too brittle and may fracture during procedure with sharp end, ideally would be cylinder, not square as sometimes swabs will be rotated to get more material. Evaluator 2: too much break point variability/polarity
HP_Spring_Polished_2 | HP_WS with hemisphere tip on a black swab w/new breakpoint | 2020-04-07 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_Spring_Polished_2.jpg | PA12 | 580 Platform | HP (LihuA Zhao; Zhao, Lihua <lihua.zhao@hp.com) |  | pass | pass
HP_WS_DG_V | HP_WS_DG with vapor polish | 2020-04-08 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_WS_DG_V.jpg | PA11 | 580 platform | HP (LihuA Zhao; Zhao, Lihua <lihua.zhao@hp.com) |  | Same design as HP_WS. Did not test | Same design as HP_WS. Did not test | pass
WIM-003 | polypropylene variant with textured tip | 2020-04-08 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wys/WIM_003.jpg | polypropylene variant |  | Richard Novak Richard.Novak@wyss.harvard.edu |  | pass | Pass (Note-npt as much cellular material as others) | pass |  | passes in terms of form factor. Might not be able to collect material. Not brittle and has appopriate strength
WIM-004 | polypropylene variant with textured tip | 2020-04-08 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Wyss/WIM_004.jpg | polpropylene variant (differs from WIM-003) |  | Richard Novak Richard.Novak@wyss.harvard.edu |  | pass | Pass | pass |  | passes in terms of form factor. Might not be able to collect material. Not brittle and has appopriate strength
MR_CLN | mass produced, sterile, flocked swab | 2020-04-08 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Miraclean/MR_CLN.jpg | ? | ? | ? |  | pass |  | pass
HP_Rev8_unpolished | HP_Rev8 no vapor polish |  |  | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_Rev8_unpolished.jpg |  |  |  |  | pass
HP_Rev8 | HP WS with central strut. Thicker neck | 2020-04-13 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_Rev8.jpg |  |  | HP (LihuA Zhao; Zhao, Lihua <lihua.zhao@hp.com) |  | Pass | Pass
HP_Rev9_Unpolished | HP Rev9 no vapor polish |  |  | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP/HP_Rev9_Unpolished.jpg |  |  |  |  | pass
HP_Rev9 | HPWS with central strut. Thinner neck | 2020-04-13 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/HP_Rev9.jpg |  |  | HP (LihuA Zhao; Zhao, Lihua <lihua.zhao@hp.com) |  | Pass | Pass
C_H_V4 | Carbon Helix V4 | 2020-04-13 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C_H_V4.jpg | Keysplint Soft Clear | Dreve PCU 90 N2 20 minutes | Carbon3D. Keshav Swarup |  | Did Not Pass |  |  | tip (and edges) are too sharp
C_H_V4S | Carbon Helix V4S | 2020-04-13 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C_H_V4S.jpg | Keysplint Soft Clear | Dreve PCU 90 N2 20 minutes | Carbon3D. Keshav Swarup |  | Did Not Pass |  |  | tip (and edges) are too sharp
C_35V3 | Carbon 35v3. Full topen | 2020-04-13 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C_35V3.jpg | Keysplint Soft Clear | Dreve PCU 90 N2 20 minutes | Carbon3D. Keshav Swarup |  | Did Not Pass |  |  | tip (and edges) are too sharp
3DPT_06_pla | white pla swab with grooved tip | 2020-04-16 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3DPrintingtech/3DPT_06_PLA.jpg | pla |  | Josh Stover. <josh@3dprintingtech.com> |  | did not pass |  | Pending | sharp tip. Can catch on anatomy when pulling out
3DPT_02_pla | white pla swab with grooved tip | 2020-04-16 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3DPrintingtech/3DPT_02_PLA.jpg | pla |  | Josh Stover. <josh@3dprintingtech.com> |  | did not pass | Did not pass. Less sharp |  | sharp tip. Can catch on anatomy when pulling out
3DPT_04_pla | white pla swab with grooved tip | 2020-04-16 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3DPrintingtech/3DPT_04_PLA.jpg | pla |  | Josh Stover. <josh@3dprintingtech.com> |  | fail-sharp tip. Can catch on anatomy when pulling out | pass. sharp
3DPT_01_pla | white pla swab with grooved tip | 2020-04-16 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3DPrintingtech/3DPT_01_PLA.jpg | pla |  | Josh Stover. <josh@3dprintingtech.com> |  | fail-sharp tip. Can catch on anatomy when pulling out | Did not pass. Sharp
3DPT_03_pla | white pla swab with grooved tip | 2020-04-16 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3DPrintingtech/3DPT_03_PLA.jpg | pla |  | Josh Stover. <josh@3dprintingtech.com> |  | fail-sharp tip. Can catch on anatomy when pulling out | pass. sharp
3DPT_01_PA12 | white PA12 swab with honeycomb tip | 2020-04-16 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3DPrintingtech/3DPT_01_PA12.jpg | pa12 |  | Josh Stover. <josh@3dprintingtech.com> |  | fail-neck is too flexible. Material is chalky and leaves residue upon handling. Tip is a little sharp, but not too bad | did not test-powder material | pass
3DPT_02_PA12 | white PA12 swab with tight honeycomb tip | 2020-04-16 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3DPrintingtech/3DPT_02_PA12.jpg | pa12 |  | Josh Stover. <josh@3dprintingtech.com> |  | fail-neck is too flexible. Material is chalky and leaves residue upon handling. Tip is too sharp | did not test-powder material
3DPT_03_PA12 | white PA12 swab with course honeycomb tip | 2020-04-16 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3DPrintingtech/3DPT_03_PA12.jpg | pa12 |  | Josh Stover. <josh@3dprintingtech.com> |  | fail-neck is too flexible. Material is is chalky and leaves residue upon handling. Tip is too sharp | did not test-powder material
3DPT_04_PA12 | white PA12 swab with honeycomb tip | 2020-04-16 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3DPrintingtech/3DPT_04_PA12.jpg | pa12 |  | Josh Stover. <josh@3dprintingtech.com> |  | fail- neck is too flexible. Matieral is chalkiy and leaves residue upon handling. Tip is the least sharp | did not test-powder material
FA_8 | Fathom spiral with 8 rotations | 2020-04-17 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Fathom/FA_8.jpg | nylon 12 |  | phone: 510-281-9000 |  | pass | did not test- residue |  |  | post autoclave: Nothing wrong post autoclave.
FA_9 | Fathom spiral with 9 rotations | 2020-04-17 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Fathom/FA_9.jpg | nylon 12 |  | phone: 510-281-9000 |  | pass | did not test- residue
FA_10 | Fathom spiral with 10 rotations | 2020-04-17 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Fathom/FA_10.jpg | nylong 12 |  | phone: 510-281-9000 |  | pass | did not test- residue
J_3_1 | resin 'flock' tip | 2020-04-20 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/OPT_industries/J_3_1.jpg | meth(acrylated urethane photopolymer |  | OPT Industries (Jifei Ou) |  | fail-I think, it doesn't bend back to shape, I thinj will have trouble getting to posterior NP. Smells Odd. | did not test- smells
C_V3_20v9.1 | carbon swab with z-axis struts | 2020-04-20 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C_v3_20V91.jpg | keysplint Soft Clear |  | Carbon 3D ( Hardik Kabaria <hardik@carbon3d.com) |  | pass. Tip seems softer, although other swabs may still be more comfortable | pass
C_Hv5_35v9 | carbon swab with sinusoidal center strut | 2020-04-20 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C_Hv5_35v9.jpg | keysplint Soft Clear |  | Carbon 3D ( Hardik Kabaria <hardik@carbon3d.com) |  | pass. Softer, narrower tip, may still not be favorite but better imo | pass
C_V3_15v1 | carbon swab with short tip | 2020-04-20 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Carbon/C_v3_15v1.jpg | keysplint Soft Clear |  | Carbon 3D ( Hardik Kabaria <hardik@carbon3d.com) |  | pass. Seems [/more]similar to envision or the others & the tip is softer | pass
PCTW | puritan sterile cotton tipped applicator w/wood handle. ref 25-826 5WC | 2020-04-20 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Puritan/PCTW.jpg | cotton and wood |  |  |  | AN-pass |  | ?
PCT | puritan cotton swab with plastic handle. ref 25-806 1PC | 2020-04-20 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Puritan/PCT.jpg | cotton and plastic |  |  |  | AN-pass |  | ?
FPR | fisher polyester-tipped applicators | 2020-04-20 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/Fisher/FPR.jpg | polyester |  |  |  | AN-pass |  | ?
3DS_NP4 | resin print tip with tight hex mesh | 2020-04-21 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3D Systems/3DS_NP2.jpg | resin |  | Patrick Dunne. Patrick.Dunne@3dsystems.com |  | Fail--tip too large. Brittle | pass
3DS_NP3 | resin print tip with tight hex mesh | 2020-04-21 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3D Systems/3DS_NP3.jpg | resin |  | Patrick Dunne. Patrick.Dunne@3dsystems.com |  | Fail--tip too large. Brittle | Did not pass
3DS_NP2 | resin print tip with tight helical mesh | 2020-04-21 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3D Systems/3DS_NP2.jpg | resin |  | Patrick Dunne. Patrick.Dunne@3dsystems.com |  | Fail--tip too large. Brittle | pass
3DS_NP1 | resin print with tight mesh | 2020-04-21 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3D Systems/3DS_NP1.jpg | resin |  | Patrick Dunne. Patrick.Dunne@3dsystems.com |  | Fail--tip too large. Brittle | pass
J_3_2 | J_3_1 with coating | 2020-04-21 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/OPT_industries/J_3_2.jpg | meth(acrylated urethane photopolymer |  | OPT Industries (Jifei Ou) |  | fail. Will crumble in anterior nares. | did not test. residue
3DPT_LN | white swab with large nubbed tip | 2020-04-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3DPrintingtech/3DPT_LN.jpg | ? |  | Josh Stover. <josh@3dprintingtech.com> |  | body is nice. Tip is too sharp and rigid | did not test. resiude
3DPT_SN | white swab with small nubbed tip | 2020-04-22 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3DPrintingtech/3DPT_SN.jpg | ? |  | Josh Stover. <josh@3dprintingtech.com> |  | Fail. TIp is too sharp and rigid | did not test. residue
J_3_3 | swab with modified bulb | 2020-04-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/OPT_industries/J_3_3.jpg | meth(acrylated urethane photopolymer |  | OPT Industries (Jifei Ou) |  | fail--too flexible and smells weird | `
3DS_NP2_1 | swab with helical tip, central strut, and rounded tip | 2020-04-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3D Systems/3DS_NP2_1.jpg | Flx-20-CLR |  | Patrick Dunne. Patrick.Dunne@3dsystems.com |  | pass if tip diameter is smaller | pass
3DS_1.5 | swab tips with a tight helical structure, with rounded tip | 2020-04-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3D Systems/3DS_1.5.jpg | Flx-20-CLR |  | Patrick Dunne. Patrick.Dunne@3dsystems.com |  | pass if tip diameter is smaller
3DS_1.9 | swab tips with a tight helical structure, with rounded tip | 2020-04-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3D Systems/3DS_1.9.jpg | Flx-20-CLR |  | Patrick Dunne. Patrick.Dunne@3dsystems.com |  | pass if tip diameter is smaller
3DS_1.7 | swab tips with a tight helical structure, with rounded tip | 2020-04-23 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3D Systems/3DS_1.7.jpg | Flx-20-CLR |  | Patrick Dunne. Patrick.Dunne@3dsystems.com |  | pass if tip diameter is smaller |  | pass
3DS_MH1_CLR | cylider for material testing | 2020-04-24 | BIDMC | https://github.com/rarnaout/Covidswab/tree/master/BIDMC/img/3D Systems/3DS_NP2_MH1_CLR.jpg | MHI-CLR |  | Patrick Dunne. Patrick.Dunne@3dsystems.com |  | N/A |  | pass
3DS_SG/MED-AMB | cylinder for material testing | 2020-04-24 | BIDMC |  | SG/MED-AMB |  | Patrick Dunne. Patrick.Dunne@3dsystems.com |  | N/A |  | *pending
C_HV5V9 | carbon swab with sinusoidal center strut | 2020-04-27 | BIDMC |  | keysplint Soft Clear |  | Carbon 3D ( Hardik Kabaria <hardik@carbon3d.com) |  | pass- a little "cheesegrater" | pass
3DS_SLS_N | nylon swab with grooved tip | 2020-04-28 | BIDMC |  | nylon-12 |  | Evan.Kuester@3dsystems.com,Zachary.Wilczynski@3DSystems.Com |  | pass--head is too large, but otherwise good | pass | pass
J_3_4 | swab with stiffer neck & bulb | 2020-04-29 | BIDMC |  |  |  | OPT Industries (Jifei Ou) |  | fail - still somewhat tacky to touch and unpleasant odor, but body design/neck strength appropriate, would try to make breakpoint more brittle | pass
3DS_NP2_FLX-20-CLR |  | 2020-05-01 | BIDMC |  |  |  | Patrick Dunne. Patrick.Dunne@3dsystems.com |  | pass - neck may still be a little to flexible, but otherwise good | pass | (pass presumably same material as 187 & 197)
3DS_NP2_MH1-CLR |  | 2020-05-01 | BIDMC |  |  |  | Patrick Dunne. Patrick.Dunne@3dsystems.com |  | fail - neck is too flexible, would likely not be able to reach posterior NP & fold in anterior nares | pass | (pass presumably same material as 187 & 197)
J_4_0 | dark gray swab with bristled tip | 2020-05-05 | BIDMC |  | meth(acrylated urethane photopolymer |  | OPT Industries (Jifei Ou) |  | Fail |  |  |  | tip is too bendy. Smells like VOCs. Tip likely to break after bending 90 degress back and forth
FL_LT_form2_hex_mtx | autoclaved clear resin with hexagonal mesh | 2020-05-06 | BIDMC |  |  |  | Sharon Soon. sharon@formlabs.com |  | fail |  |  |  | better than original, too brittle. Neck is too large (should be 3mm max) Tip is brittle and sharp
FL_LT_form2_frankenswab | autoclaved clear resin swab with rounded, nubbed tip | 2020-05-06 | BIDMC |  |  |  | Sharon Soon. sharon@formlabs.com |  | fail | pass |  |  | tip is too wide, stif
FL_LT_USF_swabs | autoclaved swabs with USF tip | 2020-05-06 | BIDMC |  |  |  | Sharon Soon. sharon@formlabs.com |  | fail | pass |  |  | better neck design than the others.Still breaks easily and has a large tip
FL_SG_frankenswabs | autoclaved SG swab with frankenswab tip | 2020-05-06 | BIDMC |  |  |  | Sharon Soon. sharon@formlabs.com |  | fail | pass |  |  | very stiff head. Other wise, form and strength of shaft is appropriate
Int_1 | autoclaved intrepid swab | [7 empty] | fail | pass | *pending |  | too brittle, but otherwise design looks appropriate
J_4_1 | black swab | 2020-05-07 | BIDMC |  | meth(acrylated urethane photopolymer |  | OPT Industries (Jifei Ou) |  | evaluator 1: pass, but a little messy. Evaluator 2: fail, material is sticky and too smelly | did not test- smelly
J_4_2 | dark gray swab | 2020-05-08 | BIDMC |  |  |  | OPT Industries (Jifei Ou) |  |  | did not test- smelly